FEIC telecom 2008-07-11

R.Hills

A.Perfetto

G.Ediss

S.Srikant

J.Crabtree

F.Schwab

T.Hunter

In the absence of DE. We went through the matrix to discuss the discrepancies for the various bands, and what needs doing to prove the results.

The following Action items were generated

- Re-analyze Band 9 data and reply to Hills questions (JC and GE)

- Send out band 3data set 95, compare x,y,z values and pol efficiency with Hills spreadsheet and NSI/JC spreadsheet (JC)

- Add better fitting (Hills spreadsheet) to our analysis to improve determination of x, y, and z (JC)

- Send out Baryshev report (GE)

- Re-analyse band 3 with cos^N with latest N values (JC and GE)

- Look at NSI software to add horn pattern correction, or do outside NSI (GE and JC), Use professional package to calculate patterns?

- Use Open ended waveguide in future measurements where S?N permits(GE)

- Do single cuts before full scans to determine setup problems in future measurements (GE and AP)

- Scans should be looked at immediately after measurement to determine if acceptable (GE and AP)

- Ask Adrian what he wants from the group and how to proceed (AP)

- Summary of meeting (GE)